PALS Coral Sea Engineering Test
and Analysis Plan

Algorithms, Diagnostic Plots, Quality Control and Decision Criteria

Persistent Autonomous LiDAR Sensor (PALS)
Engineering System Testing in the Coral Sea
Imirfada Ltd | ARIA Project
26 August 2026

Document purpose Define a reproducible engineering analysis workflow for the six-channel PALS system, with specific algorithms for waveform correction, saturation detection, overlap and alignment assessment, NF/FF stitching, polarization calibration, Raman-channel validation, motion sensitivity, uncertainty estimation, and operational quality control.

# Contents

Note: update the table of contents in Word by selecting the field and choosing Update Field.

# 1. Executive Summary

PALS is ready for engineering operation in the Coral Sea with six synchronized receiver channels: near-field co-polarized elastic, near-field cross-polarized elastic, near-field Raman, far-field co-polarized elastic, far-field cross-polarized elastic, and far-field Raman. The current engineering configuration uses fixed PMT and electronic gains; range-dependent gain ramping is not yet enabled. The campaign should therefore be treated primarily as a controlled verification of optical alignment, detector linearity, channel timing, polarization fidelity, Raman isolation, overlap behavior, and end-to-end data quality.

The analysis is designed to distinguish geometry from environment. A transmitter or receiver boresight error should affect both elastic and Raman overlap in a similar range-dependent manner. Polarization calibration errors should primarily affect the elastic co/cross relationship. Saturation should appear as loss of proportionality with pulse energy, flattened waveform peaks, or an anomalous increase in apparent depolarization when the co-polarized channel compresses. Vessel pitch and roll provide an additional natural perturbation for estimating residual alignment sensitivity.

Primary engineering decision The cruise is successful if PALS produces six stable and synchronized waveforms, the NF and FF channels can be placed on common radiometric scales, both Raman channels show physical range-resolved signals, the valid unsaturated range is quantified, residual alignment is bounded, and a defensible QC/uncertainty mask can be generated for every profile.

# 2. Objectives and Success Criteria

| Objective | Engineering question | Evidence / criterion |
| --- | --- | --- |
| Channel functionality | Are all six channels synchronized, stable and correctly identified? | Timing residuals quantified; no unexplained channel swaps; stable pre-trigger backgrounds. |
| Linearity and saturation | Are PMTs, amplifiers or ADCs compressing at fixed gain? | No clipping in interpreted bins; normalized response linear with pulse energy; linearity margin reported. |
| Alignment and overlap | Are transmitter, NF receiver and FF receiver sufficiently co-aligned? | Measured elastic and Raman overlap trends broadly match model; NF/FF ratios reach stable plateaus; weak attitude dependence. |
| Polarization | Are transmitter and receiver polarization bases stable and calibrated? | Instrument depolarization floor quantified; NF/FF depolarization agreement within uncertainty. |
| Raman performance | Are both Raman channels spectrally isolated and useful? | Repeatable water-column signal, measurable SNR, NF/FF relative scaling, no evidence that sharp elastic ghosts dominate. |
| Data quality and uncertainty | Can valid data be distinguished automatically? | Range-time QC mask and range-dependent uncertainty produced for each engineering run. |

# 3. System Configuration and Required Data

## 3.1 Optical channels

| Channel ID | Measurement | Primary engineering use |
| --- | --- | --- |
| NF_CO | Near-field co-polarized elastic | Strong-signal alignment, overlap, linearity and attenuation/backscatter proxy. |
| NF_CROSS | Near-field cross-polarized elastic | Instrument depolarization, particle-induced depolarization and saturation-bias checks. |
| NF_RAMAN | Near-field water Raman | Independent overlap/attenuation constraint and 532 nm leakage diagnostic. |
| FF_CO | Far-field co-polarized elastic | Long-range elastic response, useful range and NF/FF stitching. |
| FF_CROSS | Far-field cross-polarized elastic | Long-range depolarization and cross-talk stability. |
| FF_RAMAN | Far-field water Raman | Long-range Raman SNR, attenuation constraint and NF/FF Raman stitching. |

## 3.2 Metadata required for every acquisition block

UTC timestamp and monotonically increasing record/sequence number.

Main pulse-energy monitor, final cross-polarized dump monitor, pulse timing detector, repetition rate and laser temperature.

PMT high-voltage setpoints, fixed analog gains, amplifier configuration, digitizer input range and ADC full-scale code.

Instrument temperatures, power supply state, storage/communications state and fault flags.

Latitude, longitude, heading, pitch, roll, angular rates, heave or vertical acceleration where available, and raw inertial/NMEA record.

Operational context including platform speed, deployment geometry, daylight state, sea state, and available CTD/fluorescence/turbidity/backscatter measurements.

# 4. Processing Architecture

Processing principle Preserve Level 0 data unchanged. Every correction must generate a new variable, retain the source data, record the coefficients used, and propagate a quality flag and uncertainty contribution.

| Level | Contents | Algorithms |
| --- | --- | --- |
| L0 Raw | Original ADC arrays and metadata. | No modification. Integrity checks and checksum only. |
| L1 Engineering corrected | Time-aligned, background-corrected, energy-normalized signals. | ADC offset/polarity; trigger registration; range conversion; background estimation; fixed-gain normalization; saturation flags. |
| L2 Calibrated profiles | Range-resolved signal, SNR, overlap state and uncertainty. | Range correction; channel calibration; overlap estimates; cross-talk corrections; Raman leakage screening. |
| L3 Engineering products | Stitched profiles, depolarization, Raman constraint and QC envelope. | NF/FF scaling and blending; depolarization ratio; Raman slope; uncertainty combination; valid-data mask. |

# 5. Core Waveform Algorithms

## 5.1 Trigger registration and range conversion

The pulse timing detector defines time zero. For each channel, a fixed electronic delay is estimated from laboratory timing data or a common optical return and applied before converting sample time to range. Channel-specific residual offsets are estimated by cross-correlation around a sharp common feature.

| ri=c₀(ti-t₀-τj)/(2nw) | (1) |
| --- | --- |

Here, ri is the water-path range assigned to sample i, c₀ is the vacuum speed of light, nw is the refractive index of seawater, t₀ is the pulse trigger, and τj is the channel-specific electronic/optical delay. The factor of two accounts for outbound and return travel.

Estimate τj by maximizing normalized cross-correlation between channels over a common return feature.

Report residual alignment in samples, nanoseconds and equivalent water range.

Flag records when trigger detection is missing, ambiguous or inconsistent with expected pulse timing.

## 5.2 Background estimation and subtraction

Use one or more signal-free intervals, preferably a pre-trigger interval supplemented by a late-range interval where no physical water signal is expected. A robust estimator limits sensitivity to spikes.

| Bj=median{xji: i∈Ibg} | (2) |
| --- | --- |

| σB,j = 1.4826 × median{¦xj(i) - Bj¦ : i∈Ibg} | (3) |
| --- | --- |

| Sj(i) = sj [xj(i)-Bj] | (4) |
| --- | --- |

The MAD-based scale estimate is robust to isolated electronic spikes. The polarity factor sj  is +1 or −1, determined once for each channel. Retain both the uncorrected and corrected arrays.

## 5.3 Pulse-energy normalization

Normalize every shot or short block by the contemporaneous pulse-energy monitor. This prevents source fluctuations from being interpreted as water-column variability.

| Sjr,k= Sjr,k× Eref/EL(k) | (5) |
| --- | --- |

EL(k) is the measured pulse energy for shot k and Eref is a chosen reference, such as the median energy for the deployment. Reject or flag shots outside the calibrated energy-monitor range.

## 5.4 Fixed-gain normalization

Because gain ramping is not yet used, each deployment block must have constant PMT voltage and analog/digitizer gain. Convert digitizer counts to a channel-normalized signal using laboratory transfer coefficients.

| Sj(r)=Sj(r) / (GPMT,j Gamp,j GADC,j) | (6) |
| --- | --- |

If absolute responsivity is unavailable, retain relative calibration coefficients with traceability to the laboratory calibration file. Do not combine records acquired with different settings until the relative scaling has been verified.

## 5.5 Range correction and signal-to-noise ratio

| Xj(r)=r2Sj(r) | (7) |
| --- | --- |

| SNRj(r) =μjrσshot,j2(r) + σbg,j2 + σelec.j2 | (8) |
| --- | --- |

For engineering plots, estimate the total variance empirically from repeated shots whenever possible. Photon statistics, background variance and electronics variance should remain separable in the uncertainty budget even when their combined value is used for QC.

# 6. Saturation and Linearity Algorithms

With fixed gain, linearity must be demonstrated rather than assumed. Saturation may originate at the PMT, preamplifier or ADC. ADC clipping is obvious; PMT or amplifier compression can occur without the waveform reaching full scale.

## 6.1 ADC occupancy

| qj(r,k) =xjr,k- BjADCFS | (9) |
| --- | --- |

Flag q ≥ 0.95 as near-clipping and q ≥ 0.995 as clipped, unless laboratory characterization justifies different thresholds.

Generate time-range occupancy maps and histograms of waveform maxima.

Identify flat-topped peaks, repeated maximum codes, unusual recovery tails and baseline shifts.

## 6.2 Pulse-energy linearity test

Use natural pulse-to-pulse energy variation and, where operationally feasible, discrete reduced-energy runs. Fit the background-corrected response in each selected range interval as a function of measured pulse energy.

| Sj(r,k) = ajrEL(k) + bj(r) + εj(r,k) | (10) |
| --- | --- |

| Kj(r,k) =Sjr,kEL(k) | (11) |
| --- | --- |

In a linear chain, aj is constant and Kj is independent of EL.

Compression is indicated by declining Kj at high EL, a broken-line fit, or a nonlinear residual pattern.

Report the highest signal retaining specified linearity, plus a 6 to 10 dB operational margin.

## 6.3 Cross-channel saturation diagnostics

| QER,NF(r)=NFCOrNFRAMAN(r) | (12) |
| --- | --- |

| δNF, raw(r)=NFCROSSrNFCO(r) | (13) |
| --- | --- |

A compressed co-polarized channel causes the elastic-to-Raman ratio to fall and can bias apparent depolarization high. These ratios should therefore be plotted against both range and co-channel amplitude. Do not interpret depolarization where the co-channel is saturated or nonlinear.

# 7. Overlap and Alignment Algorithms

Near-field and far-field overlap functions depend on beam diameter/divergence, telescope aperture, field stop, relative aperture separation and optical-axis alignment. Literature on elastic/Raman lidar overlap emphasizes that misalignment changes both the onset and shape of the overlap transition, and that Raman and elastic signals can be used jointly to estimate overlap. citeturn330search13turn330search16

## 7.1 Measured overlap proxies

During intervals judged horizontally homogeneous, form pulse-energy-normalized, range-corrected signals and normalize each channel to a range where full overlap is expected.

| Ôj(r) = Xj(r) / median{Xj(r): r∈Ifull} | (14) |
| --- | --- |

This is an engineering overlap proxy rather than an absolute overlap function because water optical properties may vary with range. Use multiple intervals and both Raman channels to separate persistent geometry from environmental structure.

## 7.2 NF/FF ratio plateau

| RCOr=XNF,CO(r) / [kCO⋅XFF,CO(r)] | (15) |
| --- | --- |

| RR(r)=XNF,R(r) / [kR⋅XFF, R(r)] | (16) |
| --- | --- |

Select kCO  and kR  by robust regression in the common valid range. In the fully overlapping region, R should be close to one and approximately range-independent. A persistent slope or failure to plateau indicates relative alignment, overlap, timing or calibration error. The Raman ratio is especially valuable because it is independent of the PBS polarization split.

## 7.3 Alignment sensitivity from vessel motion

Use synchronized pitch and roll as natural angular perturbations. For homogeneous intervals and selected range bands, regress normalized signal against pitch, roll and their quadratic terms.

| Yj = β₀ + β₁ roll + β₂ pitch + β₃ roll² + β₄ pitch² + β₅ roll·pitch + ε | (17) |
| --- | --- |

A significant monotonic slope or an attitude response maximum away from the nominal orientation indicates residual boresight offset. Similar behavior in elastic and Raman channels supports a geometric cause. A change restricted to cross-polarized channels is more likely polarization-related.

## 7.4 Practical misalignment indicators

| Observation | Most likely explanation | Confirmatory check |
| --- | --- | --- |
| Elastic and Raman NF/FF transitions shift together | Transmitter/receiver or NF/FF geometry error | Attitude response and ray-trace comparison. |
| Elastic ratio changes; Raman ratio stable | Polarization or elastic-channel calibration | Inspect CO/CROSS calibration and PBS response. |
| NF and FF show opposite roll/pitch slopes | Receiver-to-receiver boresight difference | 2-D attitude response surface. |
| Ratio plateau varies with temperature | Mechanical/thermal alignment or gain drift | Correlate with bench and detector temperatures. |
| Sudden near-window increase with vessel motion | Bubbles or interface scatter | Elastic/Raman ratio, depolarization and QC camera if available. |

# 8. NF/FF Scaling and Stitching

## 8.1 Relative scale estimation

Estimate the NF-to-FF scale separately for co-polarized elastic, cross-polarized elastic and Raman families. Use only bins that are unsaturated, above the SNR threshold and within the fully overlapping common range.

| k = argmink Σr wr ρ[XNF(r) - k⋅XFF(r)] | (18) |
| --- | --- |

Here, ρ is a robust loss such as Huber loss and wr is the inverse combined variance. This prevents a small number of layers or spikes from dominating the scale estimate.

## 8.2 Blending function

| Xstitchr=wrXNFr+1-wrk⋅X_FF(r) | (19) |
| --- | --- |

Use a smooth cosine or logistic weighting function transitioning only within the verified common range. Store w(r), k, the fit residual and the uncertainty of k with the data product.

## 8.3 Stitching residual

| εstitch(r)=[XNF(r)-k⋅X_FF(r)] / {0.5⋅[XNF(r)+k⋅XFF(r)]} | (20) |
| --- | --- |

Plot the residual beneath each stitched profile. A structured residual is more diagnostic than a visually smooth blend and may indicate incomplete overlap, timing offset, misalignment, gain error or different multiple-scattering acceptance.

# 9. Polarization Calibration and Depolarization Algorithms

The receiver PBS axes define the instrument polarization reference frame. Transmitter purity is monitored with the ratio of rejected cross-polarized power to main pulse energy, while the receiver determines whether that polarization is aligned with the co-polarized basis. A Raman/polarization lidar architecture must explicitly calibrate relative channel gains and leakage; otherwise instrument cross-talk appears as environmental depolarization. citeturn330search2

| RTxk=Ecross-dumpkEmain(k) | (21) |
| --- | --- |

| δNF(r)=KNF [NFCROSS(r)-BX] / [NFCO(r)-BC] | (22) |
| --- | --- |

| δFF(r)=KFF [FFCROSS(r)-BX] / [FFCO(r)-BC] | (23) |
| --- | --- |

K_NF  and KFF are channel-balance coefficients derived from laboratory polarization sweeps. Report depolarization only where the co-polarized channel is linear and the cross-polarized channel exceeds its uncertainty-based threshold.

## 9.1 Instrument bias model

| δmeas ≈ δtrue + δinst (small-depolarization approximation) | (24) |
| --- | --- |

Estimate δ_inst during low-depolarization calibration periods and through laboratory measurements with a known linear state. For higher accuracy, use a 2×2 channel-response matrix rather than simple subtraction.

| [Cmeas, Xmeas]ᵀ = Mpol [Ctrue, Xtrue]ᵀ | (25) |
| --- | --- |

Invert Mpol only if it is well-conditioned and its uncertainty has been characterized. Compare NF and FF depolarization in the common range; a nonzero intercept in δFF versus δNF is evidence of differential instrument bias.

# 10. Raman-Channel Algorithms

## 10.1 Raman SNR and reporting range

| SNRR(r)=μRrσR(r) | (26) |
| --- | --- |

Report range limits at SNR thresholds of 1, 3 and 10. Use a contiguous-bin requirement to prevent isolated noise excursions from being labeled as valid Raman signal.

## 10.2 532 nm leakage screening

A sharp Raman-channel feature exactly coincident with a strong elastic/window return may be 532 nm leakage or ghosting rather than water Raman scattering. Screen using pulse shape, timing and the Raman-to-elastic ratio.

| LR(r)=Raman(r) / CO(r) | (27) |
| --- | --- |

Flag Raman peaks that have the same narrow impulse shape as the elastic interface return.

Compare NF and FF Raman behavior. True water Raman should follow each receiver overlap and then decay smoothly, while internal ghosts often occur at fixed instrument delays.

Confirm candidate Raman signals with the 650 nm laboratory source and filter-stack throughput test.

## 10.3 Effective Raman attenuation constraint

| YR(r)=ln{ r2PR(r) / [ELOR(r)] } | (28) |
| --- | --- |

| cR, eff=-0.5 dYR/dr | (29) |
| --- | --- |

For the engineering cruise, report this initially as an effective attenuation constraint rather than an absolute inherent optical property unless calibration, Raman cross section, wavelength-dependent attenuation and multiple-scattering effects are fully included. Use robust local regression or Savitzky-Golay smoothing before differentiation, with smoothing sensitivity included in the uncertainty.

# 11. Window Bubble and Interface-Scatter Detection

Persistent bubbles on the optical windows can dominate the near-field return, increase apparent depolarization, and reduce Raman transmission. The analysis should therefore include both prevention and automated data screening.

Pre-deployment: clean windows using the least-aggressive optics procedure; avoid residues and hydrophobic contamination; verify no haze or film under grazing illumination.

Mount/orient PALS so flow sweeps across the windows and avoid upward-facing stagnant surfaces where practical.

Conduct controlled vessel-speed or heading changes to determine whether near-window anomalies clear with flow.

Use a hydrophilic treatment only after compatibility and scatter testing on a spare coated optic.

## 11.1 Bubble QC algorithm

| Bindex = zscore(ENF,early) + zscore(δNF,early) - zscore(RamanNF,early ) | (30) |
| --- | --- |

Here, E_NF,early is integrated early-range elastic energy, δNF,early is early-range depolarization, and Raman_NF,early is the normalized Raman signal. A bubble event is expected to increase elastic return and often depolarization while reducing or not proportionally increasing Raman. Determine thresholds empirically from clean and contaminated intervals rather than treating this expression as a universal physical classifier.

# 12. Required Plot Set

| No. | Plot | Required content | Decision supported |
| --- | --- | --- | --- |
| 1 | Six-channel raw waveform panel | Six stacked panels; single shots, block averages, dashed background/noise floors, clipping level, window return and overlap zone. | Channel identity, timing, ghosts, saturation and useful range. |
| 2 | Timing and range registration | All six normalized waveforms overlaid around a common return; residual timing annotations. | Channel synchronization and range consistency. |
| 3 | ADC occupancy heatmaps | Time × range, color = fraction of full scale; 70/85/95/100% contours. | ADC saturation and near-saturation. |
| 4 | Linearity versus pulse energy | Signal and S/E versus measured pulse energy for selected range bands. | PMT/amplifier/ADC compression. |
| 5 | Cross-channel saturation ratios | CO/Raman, CROSS/Raman and CROSS/CO versus range and amplitude. | Hidden compression and depolarization bias. |
| 6 | Waveform maxima histograms | Per-channel histograms of peak counts and integrated energy. | Clipping pile-up and setting adequacy. |
| 7 | Measured overlap proxies | NF/FF co and Raman overlap proxies overlaid with modeled curves. | Geometry and misalignment. |
| 8 | NF/FF ratio versus range | CO and Raman ratios with uncertainty bands and plateau fits. | Relative scale and full-overlap range. |
| 9 | Signal versus pitch and roll | Selected range-band signal versus attitude with regression fits. | Residual boresight sensitivity. |
| 10 | 2-D attitude response surface | Roll × pitch heatmaps for NF_CO, FF_CO, NF_R and FF_R. | Direction of maximum overlap. |
| 11 | Stitched profiles | NF, scaled FF, blend, weighting function and residual for co/cross/Raman. | Stitch quality. |
| 12 | Stitching residual | Relative NF−FF residual and uncertainty across transition. | Structured calibration error. |
| 13 | NF and FF depolarization profiles | Bias-corrected δ with uncertainty, low-SNR and saturation masks. | Polarization performance. |
| 14 | NF versus FF depolarization | Scatter plot, 1:1 line, slope/intercept and range coloring. | Differential receiver bias. |
| 15 | Tx leakage versus receiver δ | R_Tx against δ_NF and δ_FF. | Separate transmitter drift from receiver/environment. |
| 16 | NF/FF Raman SNR | Signal, background, noise and SNR thresholds. | Raman useful range. |
| 17 | Raman isolation diagnostic | Raman/elastic ratio around interface and strong layers. | 532 nm leakage and ghosts. |
| 18 | NF/FF Raman consistency | Overlaid profiles, ratio and residual. | Raman scaling and alignment. |
| 19 | Raman attenuation slope | Y_R versus range, local slope, fit residual and uncertainty. | Independent attenuation constraint. |
| 20 | Time-depth curtain set | Six channels plus δ, attitude and QC masks on common axes. | System behavior through deployment. |
| 21 | Engineering stability dashboard | Pulse energy, Tx leakage, temperatures, backgrounds, channel integrals and scales. | Drift and change points. |
| 22 | Uncertainty contributions | Random, background, energy, gain, overlap, stitch, polarization and model components versus range. | NPL-style uncertainty traceability. |
| 23 | Valid-data envelope | Categorical time-range QC map. | Operational product mask. |

# 13. Daily Quick-Look Workflow

Validate file integrity, timestamps, settings and channel availability.

Generate the six-channel raw waveform panel and timing-registration plot.

Generate ADC occupancy, maxima histograms and pulse-energy linearity diagnostics.

Compute backgrounds, pulse-energy normalization, SNR and provisional usable-range masks.

Generate elastic and Raman NF/FF overlap proxies and ratio plateaus.

Generate depolarization, transmitter leakage ratio and Raman-isolation diagnostics.

Generate attitude-response plots for intervals with sufficient roll/pitch variation.

Produce one-page conclusions: valid settings, suspected saturation, alignment evidence, Raman range, bubbles/interface contamination, and recommended next run.

# 14. Uncertainty and Quality-Control Framework

For a corrected signal y=f(x), propagate first-order uncertainty using the covariance-aware GUM form. Preserve correlations for common pulse-energy normalization, shared calibration and NF/FF stitching coefficients.

| u_c²(y)=J U_x Jᵀ | (31) |
| --- | --- |

At minimum, the engineering budget should include photon/statistical variation, background estimate, energy-monitor calibration, fixed-gain calibration, overlap correction, range registration, NF/FF scaling, polarization cross-talk, Raman isolation, attitude sensitivity and model/smoothing choices.

| u_δ² ≈ (∂δ/∂X)²u_X² + (∂δ/∂C)²u_C² + 2(∂δ/∂X)(∂δ/∂C)cov(X,C) | (32) |
| --- | --- |

| u_stitch² = w²u_NF² + (1−w)²(k²u_FF² + X_FF²u_k²) + covariance terms | (33) |
| --- | --- |

UNC-ready output intent For every range-bin product retain the value, standard uncertainty, distribution/estimation method, correlation class, dependencies, and QC mask. This allows the simulated uncertainty framework to transition directly to real data and later UNC-compatible representation.

# 15. Engineering Decision Matrix

| Condition observed | Action during cruise | Post-cruise design implication |
| --- | --- | --- |
| NF_CO clipping or compression | Reduce PMT HV, analog gain, ADC range, or transmitter energy; obtain controlled energy series. | Define fixed-gain setting or implement planned gain ramp/gating. |
| Raman signal dominated by sharp elastic ghost | Check filter orientation, baffling and detector timing; obtain covered-aperture/background run. | Improve spectral isolation or stray-light control. |
| Elastic and Raman overlap both displaced | Use attitude response and alignment model to bound boresight error. | Mechanical/optical boresight correction or revised overlap calibration. |
| NF/FF ratios fail to plateau | Extend common range, verify timing/gain, use robust scaling only in valid bins. | Revisit overlap, calibration and channel geometry. |
| NF/FF depolarization disagreement | Check channel gains, PBS bias and saturation; compare with Tx leakage monitor. | Update polarization matrix and uncertainty allocation. |
| Near-window high elastic + high δ + weak Raman | Inspect/clear window; repeat at different speed/heading. | Bubble prevention and interface-scatter mitigation. |
| Stable six-channel operation with valid QC envelope | Proceed with planned science/engineering runs. | Use results to configure Phytotank calibration and real-data uncertainty validation. |

# 16. Recommended Data Product Structure

/metadata
  timestamp_utc, sequence_id, configuration_id
/navigation
  latitude, longitude, pitch, roll, heading, angular_rates, nmea_raw
/laser
  pulse_energy, cross_dump_energy, timing_pickoff, temperatures
/waveforms_raw
  NF_CO, NF_CROSS, NF_RAMAN, FF_CO, FF_CROSS, FF_RAMAN
/waveforms_corrected
  signal, background, snr, range, saturation_flag
/calibration
  timing_offsets, gain_coefficients, overlap, polarization_matrix, nf_ff_scale
/products
  stitched_co, stitched_cross, stitched_raman, depolarization, raman_constraint
/uncertainty
  standard_uncertainty, contributions, correlation_class
/quality
  valid_data_mask, bubble_flag, saturation_flag, timing_flag, low_snr_flag

# 17. References and Methodological Basis

1. Comerón, A. et al. (2023). An explicit formulation for the retrieval of the overlap function in an elastic and Raman aerosol lidar. Atmospheric Measurement Techniques, 16, 3015–3025. The paper defines the range-dependent overlap function, identifies beam/receiver geometry and relative axis alignment as key parameters, and derives an elastic/Raman overlap retrieval. citeturn330search13turn330search17

2. Wandinger, U. and Ansmann, A. (2002). Experimental determination of the lidar overlap profile with Raman lidar. Applied Optics, 41, 511–514. Cited in later overlap-method literature as a basis for combined elastic/Raman overlap determination. citeturn330search16

3. Halldórsson, T. and Langerholc, J. (1978). Geometrical form factors for the lidar function. Applied Optics, 17, 240–244. Cited by overlap literature for the geometric dependence of lidar overlap. citeturn330search16

4. Sassen, K. and Dodd, G. (1982). Lidar crossover function and misalignment effects. Applied Optics, 21, 3162–3165. Cited by overlap literature for misalignment effects. citeturn330search16

5. Rodríguez-Gómez, A. et al. (2017). An Architecture Providing Depolarization Ratio Capability for a Multi-Wavelength Raman Lidar: Implementation and First Measurements. Sensors, 17, 2957. Describes Raman/depolarization lidar architecture, calibration, stability and channel use. citeturn330search2